## Slide 1: 3D ultrasound tomography and segmentation
J. Wiskin, B. Malik, M. Lenox

## Slide 2: Quantitative 3D ultrasound tomography
Segmentation possible due to 3D quantitative accuracy
3D algorithm – necessity of 3D model of field
Proximity to the chest wall with 3D modelling of field
  Retromammary fat – glandular interface
  Retromammary fat – pectoralis interface
  “Distribution of cancers in the breast is uneven with clusters in the posterior tissue region – the retromammary fat glandular interface is an important area for cancer”

## Slide 3: Profile view – Scanner 2000 Model A
Receiver array
transmitter
Water bath

## Slide 4: QT Data Acquisition System
Receiver array
transmitter
Reflection arrays
Total field

## Slide 5: Vertical cross section of field
Vertical cross section of total field allows us to reconstruct ~30 mm above central data plane
transmitter
20 mm tall
8 row multi-element receiver
3D nonlinear full wave form inversion: Nonlinear Conjugate Gradients (RP)(Wiskin, Borup et al. 2017)
Notes: 3D nonlinear conjugate gradient minimization of ~40 million uknowns

## Slide 6: 2D vs 3D algorithm
Data levels
3D: one level of data contributes to multiple levels of the image
2D Image space
2D: one level of data affects only one level of the image
3D Image space

## Slide 7: 3D nature of transmitted field
Transmitter
Water bath
Interaction of field with Breast
receivers

## Slide 8: 3D transmitted field
Transmitter
receivers
Water bath

## Slide 9: Test phantom
holes
Low speed sphere
High speed rods
High speed sphere

## Slide 10: Reconstruction at mid-level
Low speed sphere
Low speed rod
Low speed rod
Low speed sphere
Low speed rod
Hole filled with water
High speed sphere
Water filled rods
High speed sphere

## Slide 11: Segmentation - Internal structure of acceptance phantom
Clearly shows the high speed rods (2) and high speed spheres (2) and water speed holes (4)
Segmentation based on QT Ultrasound Image
Note the ‘twist’ in the high speed rod – this objectively present – not an artefact

## Slide 12: Schematic for vertical positioning of acceptance phantom
2 mm
20 mm
Water level
5 mm cap
15 mm
receiver

## Slide 13: 15 mm spheres - ROI fast sphere
Speed of sound higher than background and water
Note the ‘twist’ in the high speed rod
Notes: Twist mentioned earlier, note the reconstruction of the low speed of sound

## Slide 14: 2D reconstruction
Sphere is not reconstructed with 2D algorithm – yellow arrow
Very top of the reconstruction – red line
Bottom of sphere
Notes: Note the bottom of sphere

## Slide 15
Near the top – sphere is visible – totally reconstructed
3D reconstruction
Top of 2D image is here
Water level
Notes: Recall the vertical extent is ~64 mm, should be roughly slightly below center of sphere: 2 mm + 5 mm = 7 mm, 7.5 mm center to top of phantom

## Slide 16: Quantitative values – can be segmented
Sphere can be segmented out in the 3D image
geometry of sphere clearly visible
segmented out from fat-like background

## Slide 17: Quantitative accuracy - Line plot through reconstruction
This level
High speed rod
Low speed sphere

## Slide 18: Line plot through 3D recon of phantom
Line plot
High speed rod
Low speed rod
QT reconstruction
True value

## Slide 19: Medium heterogeneous breast
Pectoralis visible
Inframammary fold also visible
High resolution 3D does show pectoralis
19dea
Retro-mammary fat - glandular interface
Inframammary fold

## Slide 20: Movie showing difference 2D vs 3D reconstruction
2D overlaid
Pectoralis muscle
3D image
Notes: Pectoralis visible, retromammary fat pectoralis interface retromammary fat – glandular interface

## Slide 21: 2D vs 3D Sagittal
2D overlaid on 3D
2D image
3D image
Upper extent of 2D

## Slide 22: 2D vs 3D Coronal view
2D – high resolution, high frequency image
3D – high resolution, high frequency image

## Slide 23: Scattered density breast near chest wall
Same level in red line 2D vs 3D
2D
The same level is properly reconstructed in 3D
3D extends the image at least an additional 11 mm
~30 mm of field
3D
Notes: 2D image shows degradation of resolution near the chest wall. Even before the top is reached there appears to be a significant reduction in image quality. This emphasizes the importance of 3D. Note also near the nipple region there is significant loss of resolution, again where the 3D curvature is high.

## Slide 24: Sagittal view 2D vs 3D
Resolution preserved
Additional 11 mm or more added
Red line shows same level
Note the staircase
Effect in the 2D image
Inframammary fold visible
Note the high 3D curvature

## Slide 25: 2D vs 3D – Sagittal view – small dense breast
Sagittal view – additional tissue visible – skin line visible
Pectoralis
Retromammary fat – glandular interface visible
cranial
caudal
2D reconstruction
High frequency
High definition
0.5 mm pixels

## Slide 26: 3D representation of Fibroglandular tissue - movie
Crude segmentation based on speed of sound – still of some value

## Slide 27: AI based segmentation
Based on AI: (Malik, Klock et al. 2016)
Skin, fat, connective tissue, ducts, glands, fibroadenomas and cancers

## Slide 28: Segmentation II – based on QT Ultrasound Images
Skin,
fat,
connective tissue,
ductal tissue,
glandular tissue
Movie generated from segmentation of QT Ultrasound images (DICOM)
Notes: Retromammary fat pectoralis interface, and retromammary fat glandular interface, glands pink, ducts red, cancer
Segmentation possible due to 3D quantitative accuracy
  Skin, fat, connective, ducts, glandular, cysts, fibro-adenomas, cancers
  Pectoralis is distinguishable
3D algorithm – necessity of 3D model of field
  Quantitative throughout breast volume
  Hi Resolution through breast volume: all three views
  One data level affects multiple image levels

## Slide 29: Conclusions - Quantitative 3D ultrasound tomography
Segmentation possible due to 3D quantitative accuracy
  Skin, fat, connective, ducts, glandular, cysts, fibro-adenomas, cancers effectively segmented
3D algorithm – necessity of 3D model of field
  Quantitative throughout breast volume
  Hi Resolution throughout breast volume: all three views
  One data level contributes to multiple image levels
  Pectoralis - retromammary fat interface visible

## Slide 30: QT Ultrasound, LLC
Acknowledgements
N. Pirshafiey, R. Terry, M. Racosky – design/construction of phantom
D. Borup, Ph.D. - theory
V. Zhang, Ph.D., A. Terry – discussion and experimental
J. Helms, Lightstream Animation studios – segmentation movie